## "Musts" Checklist
UV Robot "Musts" Checklist
Method of Validation
Criteria | Threshold/Must | Objective/Want | Unit | Importance | Calculation | Component Design/Spec | Comment
Physical Dimensions (WxLxH) | ≤ 0.6x0.6x2 | Accessible agile design | Meters | Gating | WxLxH | 3D Drawing or plan view | Must fit through standard bathroom stall door
Safety - Lamp Shut Off with Motion Sensing | ≤ TLV is Reached | Safe/Effective method to protect people and unit | Digital Signal | Gating | TLV by Lamp output calculation | Sensor Function | Threshold limit value (TLV) is believed to be a level to which a worker can be exposed day after day for a working lifetime without adverse effects; see attachment for incoherent UV radiation TLVs between 180 and 400nm
Safety - Steady state push force | <8kg | Stable design to avoid tipping | Steady State Force | Gating | Weight / ergonomic calculation | Total unit parts list | SEMI S8 guidelines and the estimated <8kg steady state push force is at an acceptable limit for USA/Canada ergonomic guidelines
Safety - UV On Indicator | UV "On" Light | UV "On" Light and Sound | Digital / Analog Signal | Gating | NA | Indicator / circuit | UV by definition is not visible, need indicator when the UV source is on
Automation Interface Software / Use | Basic User Interface | Tablet/Phone App | 1 or more Interfaces | High | NA | User Interface | Best if usable by by phone app
Control | Remote | Autonomous | NA | High | NA | Components to deliver functionality | Programmable for use in rooms
Cost - Total Unit BoM | ≤ $10,000 | ≤ $5,000 | US Dollars | High | US $ | By component costing | Need a low cost to allow a broader audience to purchance
Time to Kill Dose (≥25mJ/cm2) at 1.83m/6ft | ≤ 300 | Highest dose/unit time | Seconds | High | Dose calculatior UV dose, time, distance | Irradiation Validation | Dose calculation and validation.
Lamp Life Counter | Yes | Hours + On/Off Counter | "On" Hours / On/Off counter | Med | Usage time | Digital / Analog Counters | On board interface / view at a minimum.
Operation Log | Report of Time and Location | Simultaneous Localization and Mapping (SLAM) | NA | Med | NA | Accessible memory storage | Tracking of what has been cleaned for verification
Power Source | Operate on 20A 120V 50/60Hz Circuit and/or Battery | Operate on 15A 120V 50/60Hz Circuit and/or Battery | Voltage / Hz | Med | Power usage calculation | Battery Pack | Needs to operate on 80% of a 20A or a 15A 120V 50/60Hz breaker and/or use batteries
Projected Lamp Reliability | 5000 | >10,000 | Hours | Med | On/Off and supplier lifetime specs | UV Photon Source | Specify based on supplier and/or new design.
Reliability | > 6 months | 1 year | MTBF / Maintenance Schedule | Med | Quality / Prototype testing | By component reliability | Best estimate if available
Surface Proximetry Programming | ≥ 1.83 | ≥ 3 | Meters | Med | Prototype testing | Proximetry sensor | Prove functionality through a design test or prototype.
Verticle Breadth (min/max) | 0.3/2 | 0/3 | Meters | Med | Line of sight wall and floor dose |  | Height range that recieves "Time to Kill Dose (≥25mJ/cm2) at 1.83m/6ft"
Innovative Hardware Design | None | Considered and acoounted for horizontal surfaces | Feature Count | Low | NA | Feature Description(s) | Innovative designs will be scored higher, but are leaving the number of innovations
Maintenance Lamp Changes | Intuitive / safe | Can be performed by someone with limited hands-on experience in under 30min | Green to Green | Low | Estimated Time | All Lamps | Ease of lamp change will allow more users to perform
Weight | ≤ 150 | ≤ 75 | Kilograms | Low | Kg. | Robot Unit | BoM weight

## Design Req. Doc. Example
| UV Robot Design Requirements Document Example (ok to add/remove/modify requirements) | [10 empty] | Vefication Requirements
Line Item | Category | Sub-Category | Component | No. | Req No. | Title | Description | Metric | Units | Threshold (Must) | Objective (Want) | Test | Inspection | Demonstration | Analysis | Comment
1 | BATTERY TYPE | Base Battery | Li Battery | 10 | BAT10 | Base Battery Life | Device operation life | Time | hrs
2 | BATTERY TYPE | Lamp Battery | Li Battery | 20 | BAT20 | Lamp Battery Life | Device operation life | Time | hrs
3 | COMMUNICATION | WIFI | Com specs | 10 | COM10 | Communication method and specs | Details of WIFI
4 | CONTROL | Timer | Clock (Programmable) | 10 | CON10 | Programmable timer by site | Software settable | Time | Minutes
5 | CONTROL | User Interface | Phone App or Tablet | 20 | CON20 | User Control | Real-time control method | User Interface | Android/Apple
6 | CONTROL | Camera | Add to user interface | 30 | CON30 | Manual or remote operation | Remote fish eye lens | User Interface | pixels/area
7 | CURRENT | Max Current | Base Unit | 10 | CUR10 | Base current | Current peak | current | Amp
8 | CURRENT | Max Current | Lamp Unit | 20 | CUR20 | Lamp Current | Current peak | current | Amp
9 | DTC DESIGN-TO-COST |  |  | 10 | DTC10
10 | ILLUMINATION | Lamp Specs | UVC Source | 10 | ILL10 | UV Wavelength | UV Output range | Wavelength | nm +/- (x)nm
11 | ILLUMINATION | Lamp Specs | UVC Source | 20 | ILL20 | UV Source | Excimer Lamp type - serial number | Power/area | W/Cm2
12 | ILLUMINATION | Lamp Specs | UVC Source | 30 | ILL30 | Lamp length | Manufacture length | Length | cm
13 | ILLUMINATION | Lamp Specs | UVC Source | 40 | ILL40 | Number of lamps | Total lamps used | Count | numerical
14 | ILLUMINATION | Lamp Specs | UVC Source / housing | 50 | ILL50 | Lamp orientation | Installation method
15 | ILLUMINATION | Lamp Specs | UVC Source | 60 | ILL60 | Lamp output | Photon energy | Power/area | eV
16 | ILLUMINATION | Lamp Specs | UVC Source | 70 | ILL70 | Lamp lifetime | Expected/warranted lifetime | Time | Hours
17 | ILLUMINATION | Filter / Lamo Housing | Wavelength Filtered | 80 | ILL80 | Quartz/other | Description of any additional filtering needed to achieve output spectra | Wavelength | nm
18 | LOCOMOTION | Servo Motor | Base | 10 | LOC10 | Motor | Electric motor | Current | Amp
19 | MAX SPEED | Servo Motor | Base | 10 | MAX10 | Speed | Unit speed | distance/time | m/s
20 | PHYSICAL | Dimensions | Body | 10 | PHY10 | Length | Overall length | Inches | in
21 | PHYSICAL | Dimensions | Body | 20 | PHY20 | Width | Overall length | Inches | in
22 | PHYSICAL | Dimensions | Body | 30 | PHY30 | Height | Overall length | Inches | in
23 | PHYSICAL | Dimensions | Body | 40 | PHY40 | Weight | Overall length | Pounds | lbs
24 | PHYSICAL | Refector | Material | 50 | PHY50 | Material | Refectivity efficiency | transmission | eV %
25 | POSITIONING | Servo Motor | Spin on Base | 10 | POS10 | Position feedback and accuracy | Positioning accuracy | distance | cm
26 | POWER | Max Power | Base Unit | 10 | POW10 | Base Max Power | Power Rating | Power | W
27 | POWER | Max Power | Lamp Unit | 20 | POW20 | Lamp Max Power | Power Rating | Power | W
28 | POWER |  | DC/AC Inverter | 30 | POW30 | DC/AC Inverter
29 | RATED POWER | POWER | TOTAL POWER | 10 | RAT10 | System power rating | Calculation of total power needed | Power | W
30 | SAFETY | Safe Operation Training and Instruction | Declarations of Use / Warnings and Training materials | 10 | SAF10 | PPE Eyes | Personal Protective Equipment for Eyes | < TLV | mJ/cm2
31 | SAFETY | Safe Operation Training and Instruction | Declarations of Use / Warnings and Training materials | 20 | SAF20 | PPE Skin | Personal Protective Equipment for Skin | < TLV | mJ/cm3
32 | SAFETY | Audio Warning | Speaker | 30 | SAF30 | Operator warning | Operation safety to remind for PPE | Digital recording | Db
33 | SAFETY | Indicator | LED (flashing) | 40 | SAF40 | On Indicator | UV On reminder | Digital input | 0/1
34 | SAFETY | Timer | Clock (Programmable) | 50 | SAF50 | Use Counter | Software settable | Time | Minutes
35 | SENSOR | Proximetry | Break beam LED or sonar | 10 | SEN10 | Positioning sensor | Method programming position before treatment | Distance | cm
36 | SENSOR | Camera | Optional Remote operation method | 20 | SEN20 | User Control | Remote fish eye lens | Software Interface | pixels/area
37 | SENSOR | Lamp Counter | Timer for Lamp "on" | 30 | SEN30 | Software controlled | Lamp lifetime counter | Time | Hours
38 | SOFTWARE | Version | Licenced or University written | 10 | SOF10 | Onboard software | OEM or in-house software version | Build # | Build/Version #
39 | TURNING RADIUS | Servo Motor | Base | 10 | TUR10 | Navigable radius | Radius | Degrees | Deg
40 | VOLTAGE | Max Voltage | Base Unit | 10 | VOL10 | Base voltage | Voltage peak | voltage | Volt
41 | VOLTAGE | Max Voltage | Lamp Unit | 20 | VOL20 | Lamp Voltage | Voltage peak | voltage | Volt
42 | WIFI/Bluetooth | Version | Licenced or University written | 10 | WIF10 | Onboard software | OEM or in-house software version | Build # | Build/Version #